# Report 2: CGP Metrics Evaluated by CMS/Reimbursement

This report outlines the specific metrics and evidentiary standards used by CMS (specifically the MolDX program) to evaluate Comprehensive Genomic Profiling (CGP) tests for reimbursement.

## 1. Key Evaluation Metrics (The ACCE Framework)

MolDX uses the ACCE Framework (Analytical Validity, Clinical Validity, Clinical Utility, and ELSI) to determine if a test is “Reasonable and Necessary.”

### A. Analytical Validity (AV): “Does the test measure what it claims to?”

CMS requires proof that the test is accurate, reliable, and reproducible in the laboratory setting. * Key Metrics: * 1. Accuracy (Overall Concordance): “How often is the test right?” * Formula: TP+TNTotal Positions Evaluated * Note: In NGS, this is dominated by True Negatives (TN) because most of the genome is normal. It is often >99.99% and less useful than PPA/NPA. * 2. Sensitivity (≈ PPA): “If a mutation exists, do we find it?” * Formula: TPTP+FN * Numerator: Variants found by BOTH your test and the Gold Standard. * Denominator: All variants known to exist in the Gold Standard. * 3. Specificity (≈ NPA): “If it’s normal, do we call it normal?” * Formula: TNTN+FP * Numerator: Positions called “Normal” by BOTH. * Denominator: All positions known to be Normal in the Gold Standard. * 4. Limit of Detection (LoD): The lowest Allele Frequency (VAF) the test can reliably detect (e.g., 5% VAF). * Methodology (Serial Dilutions): Validated by mixing a “High Positive” sample with a “Negative/Wild-Type” sample at specific ratios. * Standard Dilution Series: 20% → 10% → 5% → 2.5% → 1.25%. * Goal: Find the lowest VAF where you still detect the variant ≥ 95% of the time (e.g., in 19 out of 20 replicates). * 5. Precision: Reproducibility (Inter-run/Intra-run).

### B. Clinical Validity (CV): “Does the result correlate with a disease state?”

CMS requires proof that the genetic variants detected are clinically meaningful. * Requirement: The test must detect variants that are: * Known “drivers” of the disease. * Predictive of therapy response (e.g., EGFR mutations for Osimertinib). * Prognostic of outcomes. * Evidence: Citations from peer-reviewed literature or guidelines (NCCN, ASCO) linking specific variants to specific cancer types. Note: “Pan-Cancer” claims require validation across multiple representative tumor types.

### C. Clinical Utility (CU): “Does using the test improve patient care?”

This is the hardest hurdle. CMS will only pay if the test result actively changes medical management. * Requirement: * The test must guide a physician to a specific therapy, clinical trial, or palliative care decision. * Evidence that testing leads to better health outcomes compared to not testing (or using standard single-gene tests). * Exclusion: “Curiosity” testing (informational only) is not covered.

### D. Case Studies: How Major Players Proved Clinical Utility & Got Paid

| Company | Strategy & Key Actions | Outcome | Code & Reimbursement (CMS 2025) | Key Reference |
| --- | --- | --- | --- | --- |
| Foundation Medicine (F1CDx) | FDA-CMS Parallel Review: First to use this pilot. Submitted one massive dossier to both agencies.Evidence: Validated on >30,000 samples demonstrating matching to therapies. | First FDA-Approved CGP.Automatic National Coverage.Defined NCD 90.2 path. | CPT 0037U~$3,500(Standard ADLT Rate) | Frampton et al., Nature Biotechnology 2013(“The validation of FoundationOne CDx”) |
| Guardant Health (Guardant360) | Liquid Biopsy Pioneer: Focused on “Tissue Issue”—proving utility when tissue biopsy fails (NSCLC).Bridging Studies: Showed high concordance with tissue to prove non-inferiority. | First FDA-Approved Liquid CDx.Secured high rate via ADLT status. | CPT 0326U / 0239U~$5,000(Initial ADLT Market Rate) | Leighl et al., Clin Cancer Res 2019(NILE Study: Non-invasive vs Invasive Lung Evaluation) |
| Caris Life Sciences | Retrospective RWE: Leveraged massive biobank to publish studies showing Overall Survival (OS) benefit for profiled patients.Tech: WES + WTS (DNA + RNA) advantage. | Coverage for WES/Transcriptome.Proved utility of RNA fusions. | PLA 0211U~$8,455(Higher rate for DNA + RNA) | Spetzler et al. (ASCO 2016)Haslem et al., Oncotarget 2016(Impact of MI on Survival in Ovarian Cancer) |

## 2. Data Requirements & Acquisition

To get paid, a lab must go through the MolDX Technical Assessment (TA) process.

### Required Data Dossier

- Validation Summary Report: A comprehensive internal document detailing the AV study (samples used, method, results).

  - Sample Size: Guidelines often suggest >59 samples representative of the tumor types and variant classes being tested.

- Sample-Level Data: Excel spreadsheets showing the raw comparison of the test vs. the reference method (Gold Standard) for every sample.

- Clinical Evidence: A bibliography or clinical study summary demonstrating CV and CU.

- SOPs: Standard Operating Procedures for the wet lab and bioinformatics pipeline.

### Acquisition Methods

- Internal Validation Study: The lab must run known samples (cell lines + clinical samples confirmed by another method) to generate the AV data.

- Clinical Studies/Trials: For Clinical Utility, labs often rely on existing literature (Level 1 Evidence) or must run their own prospective/retrospective utility study if the test is novel.

## 3. Timeline for Reimbursement

The process is lengthy and distinct from FDA approval.

| Step | Action | Estimated Time |
| --- | --- | --- |
| 1. Z-Code Application | Register test with the DEX Diagnostics Exchange to get a Z-Code identifier. | ~2 Weeks |
| 2. Technical Assessment (TA) | Submit the full dossier (AV/CV/CU data) to MolDX for review. | ~3-4 Weeks (Standard)~60 Days (Complex/Additional Info needed) |
| 3. Coverage Determination | MolDX determines if the test meets LCD (Local Coverage Determination) criteria. | Variable (Concurrent with TA) |
| 4. Pricing/Gapfill | CMS determines the payment rate (if no existing code fits). | Annual Cycle (Summer/Fall) |

Total Estimated Time: From validation completion to first paid claim ~ 3 to 6 months (best case), assuming the test falls under an existing LCD policy.

Key Note: The Z-code is mandatory. Without it, claims will be automatically rejected.

## Appendix: Key Clinical Utility References

The following papers are considered the “Foundational Evidence” that established reimbursement for modern CGP.

1. Foundation Medicine (The “Validation” Paper) * Citation: Frampton GM, Fichtenholtz A, Otto GA, et al. Development and validation of a clinical cancer genomic profiling test based on massively parallel DNA sequencing. Nature Biotechnology. 2013;31(11):1023–1031. * Significance: This paper provided the massive analytical validation (accuracy, sensitivity) data onthousands of samples that formed the backbone of their FDA submission and subsequent NCD 90.2 approval.

2. Guardant Health (The “Non-Inferiority” NILE Study) * Citation: Leighl NB, Page RD, Raymond VM, et al. Clinical Utility of Comprehensive Cell-free DNA Analysis to Identify Genomic Biomarkers in Patients with Newly Diagnosed Metastatic Non-small Cell Lung Cancer (NILE). Clinical Cancer Research. 2019;25(15):4691-4700. * Significance: Proved that liquid biopsy (Guardant360) was “non-inferior” to tissue testing and actually detected more actionable mutations (biomarkers) because it avoided the “insufficient tissue” problem. This was crucial for their ADLT reimbursement status.

3. Caris Life Sciences (The “Survival Benefit” Papers) * Citation: Haslem DS, Van Norman SB, Fulde G, et al. A retrospective analysis of predictive biomarkers in metastatic ovarian cancer with clinical outcome. Oncotarget. 2016;7(47):77235–77244. * Citation: Spetzler D, et al. Molecular profiling of 4,729 patients… ASCO Annual Meeting 2016. * Significance: These studies used retrospective data to show that patients whose treatments matched their molecular profile had significantly longer Overall Survival (OS). This “Real-World Evidence” approach allowed them to demonstrate utility for a broad WES/WTS platform without running a prospective drug trial.